## Slide 1: A CRITICAL EXAMINATION OF REGULARIZATION METHODS USED IN TWO-DIMENSIONAL AEROSOL PARTICLE SIZE INVERSION T. A. SIPKENS, S. ROGAK, MECHANICAL ENGINEERING, UNIVERSITY OF BRITISH COLUMBIA J. S. OLFERT, MECHANICAL ENGINEERING, UNIVERSITY OF ALBERTA
MOTIVATION
Proper characterization of aerosols, particularly soot, is critical to accurately modeling the anthropogenic radiative forcing responsible for climate change, to understand the processes that create these aerosols, and to appreciating their impact on human health.
Attempts to minimize the solution, that is to pick the smallest feasible solution.
0th
Working fluid
Heated saturator
Cooled condenser
Optical characterization
Aerosol Input
Mobility-selected output
Sheath flow
Electrode
Aerosol Input
Larger particles
Diffuser
Outlet filter
Inlet
Condensation particle counter
CPC
Spiral centrifuge
(similar to APM)
Differential mobility analyzer
DMA
These instruments do not measure the mass-mobility distribution directly, rather having a response that is convolved with an instrument function.
Specifically, incorporation of the instrument functions results in data represented by an integral equation:
These particles can be characterized, in part, by measuring their mass and mobility diameters. Recent studies [1] have attempted to measure the quantities simultaneously, thereby determining the mass-mobility distribution.
This can be done using a range of devices. Here, we examine data measured using a combination of a differential mobility analyzer (DMA), condensation particle counter (CPC), and aerosol particle mass analyzer (APM).
Flaring
Respiratory health
Industrial regulation
Climate change
TIKHONOV REGULARIZATION
Tikhonov regularization seeks to minimize some order derivative of the solution. In this regard, different orders of Tikhonov do different things:
Attempts to minimize the slope, that is to pick the smoothest possible solution.
1st
Attempts to minimize the curvature, that is to attempt to pick a linear solution.
2nd
This is done by adding a term to the regular least-squares formulation. Here, LTk is the Tikhonov matrix and depends on the order Tikhonov applied.
0
0.5
1.0
1.5
2.0
2.5
3.0
log10(d) [nm]
0
0.5
1.0
1.5
2.0
2.5
3.0
log10(d) [nm]
0
0.5
1.0
1.5
2.0
2.5
3.0
log10(d) [nm]
-1.0
-0.5
0
0.5
1.0
1.5
log10(m) [ag]
0th Order
λ = 2.82×10-5
0
0.2
0.4
0.6
0.8
1.0
x [counts]
×106
1st Order
λ = 5.62×10-5
2nd Order
λ = 7.08×10-5
TWOMEY AND TWOMEY-MARKOWSKI METHODS
The method proposed by Twomey and later updated by Markowski [4] is an iterative method, that is they start from an initial guess and refine it over time.
INVERSE PROBLEMS
Inverse problems occur when a quantity is derived from indirect measurement of a phenomenon.
These kinds of problems are often ill-posed, meaning that it may not be possible to find a solution unless prior information is added during the process.
To demonstrate, consider the shape of an iceberg that is floating in water.
When the shape of the iceberg is known, we can calculate what portion of the iceberg would be visible floating above water. This is the forward problem and is well-posed.
01
In contrast, when we observe an iceberg, we can often only see the portion of the iceberg that is above the water. The inverse problem is now to find the overall shape of the iceberg, given this observation. This problem is ill-posed, as there are multiple possibilities that could equally describe our observations.
02
STOCHASTIC OR BAYESIAN INVERSION
The Bayesian framework takes a statistical approach to regularizing the problem. This method is based on Bayes equation:
?
?
?
Forward problem
What would show?
Inverse
problem
What is the shape?
where N is the number of measured particles at a set point of m* and d* for the mass and mobility diameter respectively, K is the instrument function or kernel that blurs the true mass-mobility distribution, dn/dlnm·dlnd is the true mass mobility distribution, and ε represents noise in the data.
Determining the original mass-mobility distribution from the data is then an inverse problem, that is using N to find dn/dlnm·dlnd. Unfortunately, this is an ill-posed, resulting in the amplification of noise in the data that causes the inverse procedure to be challenging.
CONCLUSIONS
The present work allows for a quantitative comparison of the different methods of regularization and the use of uniform versus linear bases.
REFERENCES AND ACKNOWLEDGEMENTS
Buckley et al., J. Aerosol Sci., 114(157-168), 2017.
Voutilainen et al., Mathematical Modeling: Problems, Methods, Applications, 2001.
Carfora et al., J. Aerosol Sci., 29(10), 1998.
Markowski, Aerosol Sci. Technol., 7(127-141), 1987.
In the traditional Twomey method, the iterations take on the form to the right.
In Twomey-Markowski [4] and its adaptation to two dimensions by Buckley et al. [1], smoothing is added after applying Twomey (using a smoothing parameter, Sf). This smoothing is also applied iteratively.
0
0.5
1.0
1.5
2.0
2.5
3.0
log10(d) [nm]
0
0.5
1.0
1.5
2.0
2.5
3.0
log10(d) [nm]
0
0.5
1.0
1.5
2.0
2.5
3.0
log10(d) [nm]
-1.0
-0.5
0
0.5
1.0
1.5
log10(m) [ag]
0
0.2
0.4
0.6
0.8
1.0
x [counts]
×106
Twomey
k = 3000·Nb
Twomey-Markowski
Buckley et al., Sf = 1/100
Initial guess
k = 0
FINITE REPRESENTATION
To realize this problem numerically, the problem must be represented on some finite basis. Two main approaches exist to perform this task [3].
Simple discretization
The domain is discretized into a series of nodes at which both the kernel and true distribution are to be evaluated.
01
Finite element method
The domain is divided into a series of nodes and elements, and basis functions are defined to represent the distribution within the elements.
02
l1
l3
l2
l1
φ
l1
l3
l2
ln m, ζ
ln d, η
φ
Ωl
l4
Mesh element
Uniform basis
l3
l1
l2
l2
φ
l1
l3
l2
l1
φ
l1
l3
l2
l3
φ
l1
l3
l2
l4
φ
Linear basis
To demonstrate the ill-posedness of the problem, consider applying a simple least squared regression. The solution for the mass-mobility distribution (shown to the left) looks little like the true distribution.
Instead, prior information or regularization must be added to achieve more reasonable results.
The remainder of this work examines two aspects of this process: (i) the effect of the finite representation of these distributions and (ii) the regularization method applied [2].
Forward problem
Convolution
Inverse problem
Deconvolution
0
0.5
1.0
1.5
2.0
2.5
3.0
log10(d) [nm]
-1.0
-0.5
0
0.5
1.0
1.5
log10(m) [ag]
True mass-mobility distr.
0
0.2
0.4
0.6
0.8
1.0
x [counts]
×106
0
0.5
1.0
1.5
2.0
2.5
log10(d) [nm]
-1.0
-0.5
0
0.5
1.0
0
1
2
3
4
b [a.u.]
log10(m) [ag]
Generated data
×105
0
0.5
1.0
1.5
2.0
2.5
3.0
-1.0
-0.5
0
0.5
1.0
1.5
log10(m) [ag]
log10(d) [nm]
x [counts]
×106
0
0.2
0.4
0.6
0.8
1.0
SAMPLE PROBLEM
Log-normal distribution:
Normal distribution, effective density:
Returning to the aerosol problem, consider a bimodal distribution mass-mobility distribution. Here, each mode has the form seen on the right, where ρeff is the effective density of the particle.
For this case, the true mass-mobility distribution and the measured data will look like as follows:
Prior distribution – Adds information known a priori, that is beforehand.
Likelihood distribution – Incorporates data and information known about errors.
Posterior distribution – Combines data and prior to find a more optimal solution.
An example of one such prior is that corresponding to Tikhonov regularization, which can be rephrased in the Bayesian framework using a Gaussian form of the prior.
Another example is an exponential distance prior, which is novel to the present problem. This prior correlates pixels that are close together according to:
The matrix Ld describes the correlation length between pixels. Off-diagonal components can correlate pixels on an angle from one another. This method can be used to apply smoothing in specific directions. The results of using this prior are shown to the right.
0
0.5
1.0
1.5
2.0
2.5
3.0
log10(d) [nm]
0
0.5
1.0
1.5
2.0
2.5
3.0
log10(d) [nm]
-1.0
-0.5
0
0.5
1.0
1.5
log10(m) [ag]
Exponential
Smoothing, primary axes
Exponential
Smoothing, rotated axes
a1
a2
a1
a2
The exponential distance prior, which has not been applied to this problem previously, performs the best.
The linear basis generally performs better. However, this advantage is small.
These results give practitioners new tools to better invert aerosol data. The next step is to apply this in an experimental context to characterize soot in flares.
0
0.5
1.0
1.5
2.0
2.5
3.0
3.5
Square error, ||x – xex||2
×106
Least squares
Tikhonov (0th)
Tikhonov (1st)
Tikhonov (2nd)
Twomey (initial)
Twomey
Twomey-Markowski
Exponential
(primary axes)
Exponential
(rotated axes)
Uniform basis
Linear basis