## Slide 1: Examination of the methods available to compute the transfer function of CPMA and APM devices
Particle mass analyzers characterize aerosolized particles based on their mass-to-charge ratio. This is done using a balance of centrifugal and electrostatic forces [1,2], with the two major variants being the aerosol particle mass analyzer (APM) and centrifugal particle mass analyzer (CPMA).
MOTIVATION
Λ
1/m*, D0
m* = 0.001 fg
(D0 ≈ 0.20)
m* = 10 fg (D0 ≈ 1.2×10–4)
m* = 0.01 fg
(D0 ≈ 0.044)
CPMA | Rm = 10
0
0.2
0.4
0.6
0.8
0.1
0.3
0.5
0.7
APM | Rm = 10
0.85
0.90
0.95
1.00
1.05
1.10
1.15
m/m*
0
0.2
0.4
0.6
0.8
1.0
1.2
m* = 0.1 fg
(D0 ≈ 0.010)
m* = 0.001 fg
(D0 ≈ 0.20)
1/m*, D0
m* = 0.01 fg
(D0 ≈ 0.044)
m* = 0.1 fg
(D0 ≈ 0.010)
m* = 10 fg
(D0 ≈ 1.2×10–4)
Λ
FIG. 1 | Schematic of a particle mass analyzer
Particle mass analyzers are composed of concentric rotating cylinders with a voltage applied over the gap between the cylinders.
Characterization of these devices is given by a transfer function, which defines which particles will pass through the device at a specified setpoint.
Finite difference schemes solve this expression on a grid as a function of the axial, z, and radial, r, coordinates. This is can be done with either a Runge-Kutta or Crank-Nicholson scheme [2].
FIG. 5 | Estimated computation time
Estimates of the computational effort associated with the various methods of determining the transfer function explored in this work. Values are to be considered in relative terms only. Starred entries (*) denote those that involve a root-finding procedure to determine G0(rL).
FIG. 2
The number concentration of particles in a (left) CPMA and (right) APM. Under CPMA conditions, the particles are focused into a narrow beam; under APM conditions, the particles spread out and more particles strike the wall, even when m = m*.
FIG. 3
Transfer functions evaluated using the finite difference method across a range of setpoint masses (of a singly charged particle), holding the resolution constant.
V
Inner electrode
Outer electrode
Outlet
ω2
ω1
Bearing
Length (L)
Inlet
r1
r2
r1
r2
Inner electrode
Outer electrode
ω2
ω1
δ
rc
r
r
θ
z
vθ
vz
Insulator
FINITE DIFFERENCE APPROACH | Baseline
The transfer function can be determined if the number concentration of particles is known at the device inlet and outlet. The number concentration is described by a partial differential equation:
Convection
Diffusion
Particle migration
PARTICLE MIGRATION VELOCITY | cr
Defines the radial motion of the particle relative to the fluid and results from a balance of centrifugal and electrostatic forces
AXIAL FLOW VELOCITY | vz
Defines the distribution of the fluid velocity in the axial direction and as a function of the radial coordinate
DIFFUSION | D0
Diffusion causes a spreading of the particles throughout the device and is most relevant for small particles
Important parameters
The preferred method depends on several factors:
RECOMMENDATIONS | Which method should I use?
COMPUTATIONAL EFFORT
Optimal methods balance accuracy with the computational effort required to compute the transfer function.
REFERENCES AND ACKNOWLEDGEMENTS
1. Ehara et al., 1996. J. Aerosol Sci., 27(2).
2. Olfert and Collings, 2005. J. Aerosol Sci., 36(11).
3. Stolzenburg, 2018. Aerosol Sci. Technol., 52(10).
The APM transfer function is sensitive to the axial flow conditions, such that parabolic forms should be used when computational effort is not a limitation.
The CPMA transfer function is insensitive to the axial flow conditions, such that the plug flow assumption can be used.
For the CPMA, one should use the first-order Taylor series expansion about the center of the gap (1C) and assume a plug flow. This produces nearly identical transfer functions while improving accuracy and avoiding problems about the existence of rs.
For the APM, when diffusion is negligible, and when computational time is important, one should use the exact representation (W1) and assume a plug flow, which supersedes the previous function derived by Ehara et al. [1].
Second-order Taylor series expansions should be avoided as they are prone to numerical artifacts, such as secondary peaks, and result in little or no improvement to the overall representation of the transfer function.
T. A. SIPKENS, S. N. ROGAK | Mechanical Engineering, University of British Columbia, Vancouver, Canada J. S. OLFERT | Mechanical Engineering, University of Alberta, Alberta, Canada
CHOOSE: PARTICLE MIGRATION VELOCITY | cr
This quantity is typically written in terms of a Taylor series expansion, though several representations are available.
PARTICLE TRACKING APPROACH | New insights into a traditional approach
Particle tracking approaches seek to follow the trajectory of particles through the classifier. The techniques have been used in similar applications, e.g. by Stolzenburg [3] for the differential mobility analyzer and Ehara and coworkers [1] for the APM. We proceed by considering a slightly different framing of the particle tracking approach. The procedure is as follows:
REGION N3
Streamlines cross, unphysical
r1
r2
G0(r1)
G0(r2)
REGION F
All of the particles exit the device (Λ = 1)
REGION E
All of the particles strike the inner electrode
REGION C
Particles strike both the inner and outer electrodes
rs
r1
r2
G0(r1)
G0(r2)
REGION D
Some of the particles strike the inner electrode
rs
r1
r2
G0(r1)
G0(r2)
rs
r1
r2
G0(r1)
G0(r2)
rs
r1
r2
G0(r1)
G0(r2)
CHOOSE: AXIAL FLOW VELOCITY | vz
The axial velocity is typically considered under two conditions.
PARTICLE TRAJECTORY | F(r)
Solve an ODE to determine the trajectory of a particle given its initial radius.
MAPPING FUNCTION | G0(rL)
Evaluate the function mapping the final radial position to initial radial position
INTEGRAL BOUNDS | ra, rb
The integral bounds now depend on where the particles would theoretically originate if they were to exit at the inner and outer electrodes
EVALUATE TRANSFER FUNCTION | Λ
Use the functional form from Step 2 and the integral bounds to evaluate the transfer function
FORM OF THE TRANSFER FUNCTION
The axial velocity is used to define the functional form of the transfer function.
REGION A
All of the particles to strike the outer electrode
rs
r2
r1
G0(r2)
G0(r1)
r2
r1
G0(r1)
G0(r2)
REGION N2
Streamlines cross, unphysical
1S | Ehara et al. [1]
1st-order Taylor series expansion about the equilibrium radius, rs
1C
1st-order Taylor series expansion about the center of the gap, rc
2S
2nd-order Taylor series expansion about the equilibrium radius, rs
2C
2nd-order Taylor series expansion about the center of gap, rc
W1 | Only valid for APM
Exact expression given that ω2/ω1 = 1
GE | Most accurate
Exact expression for the more general CPMA case
PL | Plug
Uniform flow throughout the classifier
PB | Parabolic
Quadratic form based on the solution to the Navier-Stokes equation
G0(r1) < r1
Particles would have originated from inside of the inner electrode (ra = r1)
r1 < G0(r1) < r2
Particles originated within the classifier gap (ra = G0(r1))
r2
r1
G0(r1)
G0(r2)
REGION N1
Streamlines cross, unphysical
REGION B
Some of the particles strike the outer electrode
rs
r2
r1
G0(r2)
G0(r1)
r2 < G0(r1)
Particles originated outside of the outer electrode (ra = r2)
G0(r2) < r1
Particles would have originated from inside of the inner electrode (rb = r1)
r2 <G0(r2) < r1
Particles originated within the classifier gap (rb = G0(r2))
r2 < G0(r2)
Particles originated outside of the outer electrode (rb = r2)
9 CASES | 3 for G0(r1) × 3 for G0(r2)
How to find ra and rb?
INTEGRAL BOUNDS | ra, rb
Here, ra and rb collectively define the an interval of radii for particles that will exit the device. Assuming a uniform number concentration at the inlet, this interval defines
the fraction of particles that escape the classifier.
r2
r1
rb
ra
CONSIDER DIFFUSION | D0
Diffusion can be incorporated by considering particles on adjacent streamlines, here represented by a convolution with a Gaussian distribution at the inlet
4
2
3
1
5
6
7
8
0
0.2
0.4
0.6
0.8
1.0
1.2
Λ
CPMA | Rm = 10, m* = 10 fg
Case 1S ( )
Case 2S ( )
Case 2C,GE,1C ( )
Finite difference
8.5
9.0
9.5
10.0
10.5
11.0
11.5
m [fg]
(a)
0
0.2
0.4
0.6
0.8
0.1
0.3
0.5
0.7
APM | Rm = 10, m* = 10 fg
Case 1S ( )
Case 2S ( )
Case 1C,2C,
W1,GE ( )
Finite difference
8.5
9.0
9.5
10.0
10.5
11.0
11.5
m [fg]
(b)
-0.06
-0.04
-0.02
0
0.02
-0.1
0
0.1
0.2
ΔΛ
FIG 4 | Non diffusing, plug flow
0
0.2
0.4
0.6
0.8
0.1
0.3
0.5
0.7
8.5
9.0
9.5
10.0
10.5
11.0
11.5
Case 1C,W1 ( )
Case 1S ( )
-0.06
-0.04
-0.02
0
0.02
Finite difference
APM | Rm = 10,
m* = 10 fg
0
0.2
0.4
0.6
0.8
1.0
1.2
Λ
CPMA | Rm = 10, m* = 10 fg
Case 1S ( )
Case 1C ( )
Finite difference
8.5
9.0
9.5
10.0
10.5
11.0
11.5
-0.1
0
0.1
0.2
ΔΛ
FIG 5 | Non-diffusing, parabolic flow
0
0.2
0.4
0.6
0.8
1.0
1.2
Λ
Case 1S ( )
Case 2S ( )
0.85
0.90
0.95
1.00
1.05
1.10
1.15
-0.1
0
0.1
0.2
ΔΛ
Finite difference
Case 1C,2C ( , )
CPMA | Rm = 10,
m* = 0.01 fg
0.85
0.90
0.95
1.00
1.05
1.10
1.15
0
0.2
0.4
0.6
0.8
0.1
0.3
0.5
0.7
Λ
Case 1S ( )
Case 2S ( )
Case 1C,2C,W1 ( )
Finite difference
-0.04
-0.02
0
0.02
APM | Rm = 10,
m* = 0.01 fg
ΔΛ
FIG 6 | Diffusing, plug flow
m [fg]
(a)
m [fg]
(b)
m/m*
(a)
m/m*
(b)
1S
1C
2S
2C
GE*
1S*
1C*
1S
1C
2S
Tri.
2C
PL
PB
PL
Non-diffusing
Diffusing
CPMA | Rm = 10, m* = 0.01 fg
FD
1S
1C
2S
2C
GE*
1S*
1C*
1S
1C
2S
Tri.
2C
W1
W1*
W1
PL
PB
PL
Non-diffusing
Diffusing
APM | Rm = 10, m* = 0.01 fg
FD
10–3
10–2
10–1
100
101
Estimated CPU time [s]
m = 0.9m*
m = m* = 0.01 fg
m = 1.1m*
Outer electrode
Inner electrode
r1
r2
r
r1
r2
r
r1
r2
r
0
0.2
0.4
0.6
0.8
1.0
1.2
n/n0
L
0
z
L
0
z
Inlet
Outlet
Inlet
Outlet
APM | Rm = 10
CPMA | Rm = 10
Inner electrode
Outer electrode